# Multi-User Pump Assessment Chart

## Clinic Location 		Elevation

Standard Values for Medela Pumps

Symphony	50 – 250 mmHg

## Lactina

- MIN setting	90 – 100 mmHg

- MAX setting	240 – 250 mmHg

(Values should go up as dial is adjusted from MIN to MAX)

## Pump in Style

First phase	         50 – 200 mmHg Second phase          100 – 250 mmHg

- Low setting      90 – 100 mmHg

- High setting    240 – 250 mmHg

Standard Values for Hygeia Pumps

EnDeare

- MAX setting		230 mmHg or greater

(suction is low if less than 230 mmHg)

EnJoye

- MAX setting		230 mmHg or greater

(suction is low if less than 230 mmHg)

| Date | Serial Number (if applicable) | Value on MIN/Low | Value on MAX/High | Is pump function normal? Yes No | Comments | Name of Tester |
| --- | --- | --- | --- | --- | --- | --- |

Multi-User Pump Assessment Chart        					Page 2

| Date | Serial Number (if applicable) | Value on MIN/Low | Value on MAX/High | Is pump function normal? Yes No | Comments | Name of Tester |
| --- | --- | --- | --- | --- | --- | --- |